## Slide 1: Inertial Measurement Analysis
OCE2901 – Surf Engineering Analysis

## Slide 2: Introduction
What are IMUs?

## Slide 3: Inertial Measurement Units
Come in a variety of form factors from massive mechanical meters to microscopic moving mechanisms
Different grades:
Consumer
(Cell phones)
Automotive
(Self-driving)
Industrial
(Research)
Military
(Missiles, aircraft)
Navigational
(Submarines)
Cost, precision, and accuracy (exponential)

## Slide 4: Degrees of Freedom (DOF)
IMUs are defined by their sensors and DOFs
Each measurement axis counts as one DOF
DOFs | Accelerometer | Gyroscope | Magnetometer | Barometer
3-DOF | 3-axis | 3-axis | - | -
6-DOF | 3-axis | 3-axis | - | -
9-DOF | 3-axis | 3-axis | 3-axis | -
10-DOF | 3-axis | 3-axis | 3-axis | 1-axis
OR

## Slide 5: Accelerometers
Measure change in velocity over time (acceleration)
Why is the Z-axis not reading 0?
Does this only occur with the Z-axis specifically?
Offsets due to small errors in manufacturing or temperature drift can cause the accelerometer to read non-zero values when stationary
Bias caused by gravity or misalignment causes a tilt towards a specific direction and report incorrect acceleration
The displacement of a small mass due to a force can be correlated to an acceleration on a scale

## Slide 6: Calibrating Accelerometers
Calibration is essential to reduce errors from offsets and bias
1. Take measurements in different orientations
Pitch up, pitch down, roll left, roll right, etc.
2. Record accelerations in a data file
Tab- or comma-delimited
3. Use a software program to fit the data to an ellipsoid
The ellipsoid should be centered about the origin (XY) or 1g (Z)
4. Visualize the results
The calibrated ellipsoid should be roughly spherical
Apply biases and offsets to accelerometer and recollect the data
Source: https://github.com/michaelwro/accelerometer-calibration,
https://www.youtube.com/watch?v=-1tmYPE7MAQ

## Slide 7
Calibrating Accelerometers (cont.)

## Slide 8: Gyroscopes
Measure the angular rotation rate of a body through the Coriolis effect
A forced oscillation is induced on the x-axis. While in motion, if an angular velocity, , is induced on the z-axis, a force will be imparted on the y-axis and therefore be displaced proportionally.
Over time, the gyroscope will drift due to procession and the accumulation of minute errors. This will cause non-zero readings, even when stationary
Proof in the Appendix and lecture notes

## Slide 9: Calibrating Gyroscopes
Temperature compensation
Use an algorithm to correct the gyroscope sensitivity based on temperature
Bias correction
Measure the gyroscope outputs in the steady state condition and subtract this offset from the readings
Scale factor adjustment
Adjust how sensitive the gyroscope is to angular rotation
Kalman Filtering
We will discuss later

## Slide 10: Magnetometers
Measure intensity and direction of magnetic fields using magnetoresistive elements
What does “magnetoresistive” mean?
Atoms within a magnetoresistive element change their orientation with the magnetic field which can hinder the path of free electrons (current).
Measuring the resistance allows us to measure the magnetic field.
Offsets due to small manufacturing defects, temperature drift, hard iron and/or, soft iron interference can cause the magnetometer to read incorrect values
Bias caused by magnetic declination or misalignment can cause the magnetometer to skew towards a certain heading (compass error)
What is “magnetic declination”?
Current
Current

## Slide 11: Calibrating Magnetometers
Hard Iron Distortion
Interference by a permanent magnetic field
Offset_x = (max(x) + min(x)) / 2
Offset_y = (max(y) + min(y)) / 2
Offset_z = (max(z) + min(z)) / 2
Corrected_x = sensor_x – offset_x
Corrected_y = sensor_y - offset_y
Corrected_x = sensor_z – offset_z
Soft Iron Distortion
Interference by materials or fields that distort magnetic fields
Avg_dx = (max(x) + min(x)) / 2
…
Avg_d = (avg_dx + avg_dy + avg_dz) / 3
Scale_x = avg_d / avg_dx
…
corrected_x = (sensor_x - offset_x) * scale_x
Source: https://www.appelsiini.net/2018/calibrate-magnetometer/

## Slide 12: Calibrating Magnetometers (cont.)
Hard Iron interference caused by a permanent magnetic next to sensor
Hard Iron interference caused by a speaker next to sensor after magnet is removed
Note: tenfold reduction in scale from previous plot

## Slide 13: Calibrating Magnetometers (cont.)
New plot showing soft iron interference after the hard iron interference is compensated
Final plot with fully compensated magnetometer
Why might these two plots not look so different?
Does the magnetic calibration depend on where you are? How?
Why is this not the case for the accelerometer or gyroscope?

## Slide 14: Demonstration!

## Slide 15: IMU Processing
Reading in the data like eye-em-you

## Slide 16: Overall Process
Record
Use the IMU and controller to record data
Typ. CSV file or memory structure
Pre-process
Filtering
Decimation
Extract
Features
Windowing
Statistics
Normalization
Spectral Analysis
Analyze
Use the filtered data such that you can answer your research question
Calibrate
Put the sensors into known positions to calculate offsets/biases

## Slide 17: Calibration
These sensors have biases and scale factors that account for the sensitivity, misalignment, drift, and error in the measurements
Calibration is the process of calculating these values to improve the accuracy of the sensor
Regular calibration ensures that your data is accurate and the IMU is functioning properly

## Slide 18: Pitch Calibration

## Slide 19: Roll Calibration

## Slide 20
Pitch Calibration
Stationary, lying flat
Roll Calibration
Stationary, lying flat
Stationary, lying flat
On Tail
Flat
On Nose
Port
Flat
Starboard

## Slide 21: Recording
IMU Data Structure
Acceleration X,
Acceleration X, 𝑎_𝑥
Acceleration Y,
Acceleration Y, 𝑎_𝑦
Acceleration Z,
Acceleration Z, 𝑎_𝑧
Roll Rate X,
Roll Rate X, 𝜔_𝑥
Roll Rate Y,
Roll Rate Y, 𝜔_𝑦
Roll Rate Z,
Roll Rate Z, 𝜔_𝑧
Mag Intensity X,
Mag Intensity X, 𝐵_𝑥
Mag Intensity Y,
Mag Intensity Y, 𝐵_𝑦
Mag Intensity Z,
Mag Intensity Z, 𝐵_𝑧
IMU Data Structure
, | , | , | , | , | , | , | , | 
0.02, | 0.04, | 9.87, | 0.15, | 0.46, | 0.47, | 0.37, | 0.37, | 0.04
0.00, | 0.05, | 9.90, | 0.22, | 0.43, | 0.49, | 0.38, | 0.37, | 0.05
0.00, | 0.06, | 9.91, | 0.19, | 0.44, | 0.51, | 0.38, | 0.37, | 0.05
0.01, | 0.11, | 9.88, | 0.16, | 0.45, | 0.52 | 0.38, | 0.37, | 0.05
0.02, | 0.04, | 9.87, | 0.15, | 0.46, | 0.47, | 0.37, | 0.37, | 0.04
0.00, | 0.05, | 9.90, | 0.22, | 0.43, | 0.49, | 0.38, | 0.37, | 0.05
0.00, | 0.06, | 9.91, | 0.19, | 0.44, | 0.51, | 0.38, | 0.37, | 0.05
0.01, | 0.11, | 9.88, | 0.16, | 0.45, | 0.52 | 0.38, | 0.37, | 0.05
CSV Log File
IMU

## Slide 22
Start
Function: Record IMU Data
IMU Sample Interval
Has the IMU Sample Interval Passed?
Sample IMU Values and Save to Memory
IMU Telemetry Structure
(Optional) Update Sensor Fusion Model
Filename
Save Telemetry Structure to Encoded Log File
Decode Log File for Processing
Return CSV File of Logged Data
No
Yes
Initialize/Update Time of Last Log
Continue
Post-Deployment

## Slide 23: Pre-Processing
After we offload the data from the instrument, we can begin treating it to make our analysis more accurate
We can accomplish this using three different techniques:
Filtering
Decimation
Normalization

## Slide 24: Filtering
Low-Pass
High-Pass
Band-Pass
Band-Stop
We can use different filtering techniques to attenuate unwanted noise in our data at the cost of overall accuracy and a phase shift
What are the different types of IIR filters we can use?
What other filter types exist?

## Slide 25: Decimation
Decimation is the process of removing excess data points as a part of filtering, or to reduce the computing costs.
Commonly, decimation occurs by taking every Nth sample, or averaging ever N samples (moving average filter)
EX: If a sensor samples a 16 Hz signal at 1 MHz, then it will be more computationally efficient to decimate the raw signal to 64-100 Hz

## Slide 26: Normalization
Normalization refers to the process of scaling a dataset down to a “normal” scale
Typically, this means the mean of the data will be zero and the data points are shrunk to +/- 1 standard deviation
EX: Two sensors of different sensitivities measure the same signal and are normalized to +/- 1g of acceleration for comparison

## Slide 27: Feature Extraction
How can we determine what data are important?

## Slide 28: Windowing
Windowing functions break the data down into specific windows that are useful for reducing high frequency noise and eliminating artifacts when overlapped and recombined
There are multiple functions that accomplish different tasks:

## Slide 31: Statistics
Mean
Median
Mode
Range
Correlation
Uncertainty
The average of all the values in the time series
The middle value of the time series
The most recurrent value in the time series
The extent of the time series’ maximum and minimum
How two different time series relate to each other
Errors or bias associated with the time series that decrease accuracy

## Slide 32: Spectral Analysis
The time series is converted to the frequency space via a Fourier Transform and key features can be identified
What do you notice about this graph?

## Slide 33: Analysis
You can apply any of the techniques that you need to in order to answer your research question
If you need velocity and position, you may want to consider integrating over a specific window to reduce error
If you want energy, you will need to use kinematics, or convert the acceleration signal in the board to the frequency space and look at the power density
One thing that will be useful is converting both the surfer motion and wave motion into the same reference frame – we can do this by finding the surfer’s orientation!

## Slide 34: Attitude and Heading Reference System (AHRS)
An AHRS is a 9DOF IMU equipped with a triad of accelerometers, gyroscopes, and magnetometers in all three axes
Sensor data are fused together to estimate a body’s inertial orientation in 3D space
This Photo by Unknown Author is licensed under CC BY-SA-NC

## Slide 35: AHRS Design
Goal of an AHRS is to fuse multiple data streams to produce an estimate of the body’s orientation and heading in 3D space (pose)
IMU
AHRS
GPS
Mathematical Models
Pitch,
Yaw,
Roll,
What is an alternative output vector?

## Slide 36: AHRS Design
Goal of an AHRS is to fuse multiple data streams to produce an estimate of the body’s orientation and heading in 3D space (pose)
IMU
AHRS
GPS
Mathematical Models
Quaternion Vector

## Slide 37: Why Does This Matter?
Let’s consider a ship at sea:
Can we evaluate the ship’s orientation?
N
If we want to ship to go along a course, all we need to do is just follow the compass for a certain time and speed!
Dead-reckoning navigation

## Slide 38: Why Does This Matter?
Now the ship encounters waves directly front:
What happens to a tilted compass? Why?
N
A tilted compass will sense the magnetic field differently and therefore be inaccurate!
Therefore, we can no long rely on dead-reckoning for navigation
How was this mitigated in the past? (Hint: the specific artifact is on the Link second floor)

## Slide 39: Why Does This Matter?
Now the waves are approaching from the bow-portside quarter:
What will the vessel want to do?
N

## Slide 40: Why Does This Matter?
Now the waves are approaching from the bow-portside quarter:
N
The vessel will want to turn side-on to the waves and begin to roll. Now we have an unknowable uncompensatable, force acting on the vessel and we are going all over the place.
Dead-reckoning is just that – DEAD!

## Slide 41: In real-world scenarios, knowing the orientation of our vessel is extremely important as it allows us to better navigate and understand our environment
Why Does This Matter?

## Slide 42: Rotation and Reference Frames

## Slide 43: Orientation
A body can be rotated in 3D space along the X, Y, and Z axes
  Roll (): Rotation about the x-axis
  Pitch (): Rotation about the y-axis
  Yaw (): Rotation about the z-axis
These are known as the Euler Angles and are easy for us to understand conceptually

## Slide 44: Rotation
Rotation is how a body rotates in 3D space relative to a central coordinate system
We can define this as Euler angles (roll, pitch, yaw)
Are we rotating now? In what regards?

## Slide 45
Z
Y
X
Local Coordinate System
X-Y-Z
The local inertial reference frame used to represent rotations and motions relative to an arbitrary body

## Slide 46
N
E
D
Universal Coordinate System
North-East-Down
The global inertial reference frame used to represent motions and rotations relative to the center of the Earth*
(Assumes N-E plane of motion is tangential to surface)

## Slide 47
N
E
D
Reference Frame Rotations
We can place a body with its own local coordinate system onto the global coordinate system.
But, how can we relate the two? I.e. how can we determine what “acceleration in the X-direction” means in terms of “acceleration in the East-direction”?
Z
Y
X

## Slide 48
N
E
D
Reference Frame Rotations
We can rotate the body frame of reference to the global frame and directly correlate each value in our data vector. E.g.
We can accomplish this using a Rotation Matrix
Z
Y
X

## Slide 49: Rotation Matrix
A rotation matrix is a mathematical model for translating one body’s inertial reference to another
  E.g., the local body to the global body
We can use Eulerian transformations of the vectors
  I.e., you will need to know the roll, pitch, and yaw rotations
Let’s start with a 2D rotation matrix:
Note: square matrix (orthogonal determinant = 1)

## Slide 50: Rotation Matrix
If we want to rotate a vector , degrees, then we must do the following:
Imagine we have a vector, , in the X-Y plane with is point defined at as shown to the left
We can express in polar form as:

## Slide 51: Rotation Matrix
We can rotate about the Z-axis by degrees to obtain a new vector, :
The new vector will have the coordinates
We can express in polar form as:

## Slide 52: Rotation Matrix
By expanding and and using trigonometric identities, we obtain:

## Slide 53: Rotation Matrix
We can then express the resultant equations, and , in the form of a 2x2 rotation matrix:

## Slide 54: Rotation Matrix Example
We can use this rotation matrix to determine how much the current from the ENE is hindering our vessel:
X
Y
D
N
E

## Slide 55: Rotation Matrix Example
We can use this rotation matrix to determine how much the current from the ENE is hindering our vessel:
D
N
E
X
Y
Where is the current velocity apparent to the ship’s local reference frame

## Slide 56: Rotation Matrix Example
We can also apply this same principle to determining our acceleration/velocity/displacement in the global frame
X
Y
N

## Slide 57: 3D Rotation Transformations
In 3D space, rotation can be performed about the x-, y-, or z-axis.
A basic rotation about a single axis is called an “elementary rotation” and can be represented by the following:

## Slide 58: 3D Rotation Transformations
To rotate a 3D vector, select the order and axes which you want to use and multiply their rotation matrices and the vector together.
  E.g. to rotate a body’s XYZ accelerations into the NED reference frame, you could do the following
This rotates the body acceleration first from roll, then pitch, then yaw, to the NED reference frame
Why is this not the ideal solution?

## Slide 60: Gimbal Lock!
When two axes become aligned, changes to either axis become irrelevant – thus the system loses one degree of movement
In gimbal lock cases rotation and plotting yield discontinuities which are difficult to manage
Solution 1: Update your rotation order when you detect gimbal lock, or drive two axes simultaneously
What might be a different solution? Hint: You’ll need to add a fourth (quat) dimension.

## Slide 62: What is a Quaternion?
A quaternion is a complex number representing an orientation in 3D space
It can describe the rotation of a point about an axis without the limitations of Euler singularities (gimbal lock)
Suggested sources:
  J. B. Kuipers. Quaternions and Rotation Sequences: A Primer with Applications to Orbits, Aerospace and Virtual Reality. Princeton University Press, 1999.
  https://www.youtube.com/watch?v=d4EgbgTm0Bg
  https://eater.net/quaternions/

## Slide 63: Quick Quaternion Derivation
Complex numbers in form a plane and a relationship with 2D geometry via the unit circle
where
This can yield a rotation:
Where the vector is rotated about the z-axis by degrees

## Slide 64: Quick Quaternion Derivation (cont)
In , we can define any complex number as:
Additionally, any complex number will have a length (magnitude) as defined by the Pythagorean Theorem:
We can also add/subtract in :
And we can multiply:

## Slide 65: Quick Quaternion Derivation (cont)
Euler states that exists in a unit circle on the plane.
Therefore, we can represent complex numbers in polar coordinates where the magnitude of the number is multiplied by where is the rotation about the z-axis. We can redefine:
and
Then, when we multiply them together:
In short, to multiply two complex numbers, we multiply their lengths and add together their rotations about the z-axis to get the resultant vector

## Slide 66: Quick Quaternion Derivation (cont)
So, if we take a complex vector, , and multiply it by a unit complex vector, , we have a resultant rotation:
This new complex vector has the same magnitude as , but is rotated degrees
Therefore, we can use complex arithmetic to accomplish a geometric operation!

## Slide 67: Quick Quaternion Derivation (cont)
It is impossible to achieve the same in a 3D complex space because of a theorem that supposes you cannot multiply complex triplets.
However, Hamilton (not the Founding Father) determined that you can multiply complex quadruplets to represent a 3D rotation.
This was the genesis of the quaternion as defined by:
Does this look familiar to anyone who remembers basics physics? Why?

## Slide 68: Quick Quaternion Derivation (cont)
Quaternions are a 4D vector that we can say have a scalar and a 3D vector part:
We can use this property to perform dot and cross products with other quaternions (Hamiltonion products)

## Slide 69: 3D Rotation Using Quaternions
3D rotation can be accomplished by multiplying a point vector by a unit quaternion and its conjugation
Where is the original point vector,
is the unit quaternion
is the conjugation of the unit quaternion
This means we can take any known vector and rotate it about an axis without risking gimbal lock!
Note:

## Slide 70: Linear Acceleration Example
N
E
D
X
Y
Z
G
In this case, let’s assume a surfer is on the water with waves approaching from their port-aft quadrant. They will begin to yaw, pitch, and roll simultaneously as shown to the left. At any given moment, gravity has an influence on all three axes of measurement for the accelerometer. How do we determine the effects of gravity on all three axes?
First, we know that the board has a certain orientation (roll, pitch, yaw) from flat (X-Y N-E) and pointing north (). We can express this as a quaternion that is our body frame orientation () relative to the global frame () denoted as, .
By determining this quaternion (more on that later), we can calculate the rotation of gravitational acceleration from the Global () to the body () frame:
We can then subtract from our measurements to determine the linear acceleration experienced by the surfboard in the local frame

## Slide 71: Converting Between Quaternion and Euler
In order to convert a quaternion to an Euler representation we must account for
  Euler rotation matrices
  The Eulerian sequency of rotation
If we want to perform a Yaw-Pitch-Roll (3-2-1) rotation, recall:

## Slide 72: Converting Between Quaternion and Euler
Through mathematical black magic, we can create a relationship from these matrices such that:
Note: This conversion only applies for the 3-2-1 rotation. Other definitions exists for other rotation sequences.

## Slide 73: Converting Between Quaternion and Euler
To go from quaternions to Euler, the process is simpler and does not depend on the rotation sequence:

## Slide 74: Sensor Fusion

## Slide 75: Introduction
The fundamental goal of sensor fusion is to take data from a variety of sources and mix it all together to accomplish an estimate of the object’s state at that point in time.
IMU
AHRS
GPS
Mathematical Models
Quaternion Vector

## Slide 76: Fast Complimentary Filter
A complimentary filter fuses information together using basic physical relationships
Source: Wu, Jin & Zhou, Zebo & Chen, Jingjun & Fourati, Hassen & Li, R.. (2016). Fast Complementary Filter for Attitude Estimation Using Low-Cost MARG Sensors. IEEE Sensors Journal. 16. 1-1. 10.1109/JSEN.2016.2589660.
Gyroscope Orientation,
Accel/Mag Orientation,
Why should we weight each orientation estimation? What will happen to the gyro over time?

## Slide 77: FCF – Overview
Accelerometer
Gyroscope
Magnetometer
1st Complimentary Filter for Gravity Estimation
2nd Complimentary Filter for
Grav-Mag Fusion
Magnetic Distortion Present?
Roll, Pitch Estimate ONLY
Complete Body Pose Estimate
Yes
No

## Slide 78: FCF – Gyroscope Orientation
We can define the angular rate as the change in orientation over time, . We can then describe the resultant relationship:
Where is the body pose as determined from the gyroscope at timestep, , and is the update time of the filter

## Slide 79: FCF – Accelerometer Orientation
Errors with the gyroscope (misalignment, noise) causes a drift within the integration over time.
We can correct it by updating our pose using an accelerometer which is more stable over time

## Slide 80: FCF – 1st Complimentary Filter
We can now represent the body pose recursively with:
Type equation here.

## Slide 81: FCF – 1st Complimentary Filter
We can combine the two estimates from the gyroscope and accelerometer together to get an initial estimate of our orientation:
The hyper parameter, , determines which guess is preferred
What does this orientation estimate tell us? Is it a complete estimate?
How would you tune if you had a noisy/unreliable gyroscope? What if it was accurate initially, but drifted?

## Slide 82: Appendix

## Slide 83: Accelerometer Calibration Psuedocode

## Slide 84: Accelerometer Calibration Flow Chart

## Slide 85: Gyroscope Proof (single axis)
The position of the center of mass in the body frame is represented by:
  (1)
The inertial velocity of the mass is given by the positional derivative and the tangential velocity due to rotation:
  (2)
The inertial acceleration is the next derivative given by:
  (3)
The first element in (3) represents the acceleration experienced by the driven axis (x-axis); the second element is that of the sensing axis (y-axis).
Recall Newton’s Second Law, . If we start from rest, therefore,
  (4)

## Slide 86: Gyroscope Calibration Pseudocode

## Slide 87: Gyroscope Calibration Flow Chart

## Slide 88: Magnetometer Calibration Pseudocode

## Slide 89: Magnetometer Calibration Flow Chart